# Feasibility of metabolic imaging of hyperpolarized 13C-pyruvate in human breast cancer

# M. A. McLean, C. J. Daniels, J. T. Grist, R. F. Schulte, T. Lanz, A. Chhabra, H. Earl, B. Basu, I. Wilkinson, D. J. Lomas, C. Caldas, J. Abraham, M. J. Graves, F. J. Gilbert, K. M. Brindle, F. A. Gallagher

# Introduction

. Imaging of the breast with hyperpolarized 13C yields new challenges compared to imaging the prostate [1]. E.g. large anteroposterior B0 gradients [2] require correction and the anatomy and patient positioning need a new, highly optimized RF coil array for achieving sufficient SNR/spatial resolution. As a first step, we have investigated single-breast imaging in the coronal plane.

Methods

A BRCA gene carrier with a 38-mm diameter grade 3 triple-negative invasive ductal carcinoma was studied on a 3T MRI (GE Healthcare) using a prototype 8-channel 13C breast coil (Rapid Biomedical), containing 2 transmit/receive coils and 6 receive-only covering both breasts in a prone position. 1H imaging was performed with the body coil. Following injection of 40ml of 250mM 13C-pyruvate, polarized to c. 25%, a 1-minute time series of spirals with IDEAL encoding (3) was collected (flip angle 10°, TR=260ms, 8-step cycle, time resolution 2.08s, 3 x 3-cm thick slices, 3mm gap, 40-pt spiral, 24cm coronal FOV, real pixel size 12 x 12 x 30mm). IDEAL reconstruction of images was optimized separately for each slice to enable independent frequency offsets to be applied. Kinetic modelling was performed in MATLAB, with automated tumour segmentation.

Results

Tumour pixels were identified by the segmentation algorithm only in the tumour-containing slice 2, and the average estimated flux from pyruvate to lactate kPL within this ROI was 0.022 s-1 (Fig. 1). The frequency shift of pyruvate relative to slice 2 was +6 Hz in slice 3 and -34 Hz in slice 1, confirming a sharp gradient in B0 approaching the nipple, which was corrected by optimizing slices separately (Fig 2). Images of lactate and pyruvate summed over the time course (Fig 3) showed strong signal of both metabolites over the tumour in slice 2, lower pyruvate in the slice toward the chest wall, and no consistent signal in slice 1.

Conclusion

This first-in-Europe study in breast cancer established the feasibility of obtaining metabolite images with high temporal and moderate spatial resolution in humans in vivo following administration of hyperpolarized 13C-pyruvate. Coronal image orientation allowed application of significant corrections for a known limitation, the anteroposterior B0 gradient, as well as a small FOV to improve spatial resolution. Kinetic rate constants within the tumour were found to be consistent with previous reports in human prostate cancer (1).

References

1) Nelson SJ et al. Sci Transl Med 5, 198ra108 (2013). 2) Maril N, et al. Magn. Reson. Med. 2005; 54:1139-1145. 3) Wiesinger F, et al. Magn Reson Med 2012; 68:8-16.

Figure 1: (Left) Axial scout image showing positions of coronal slice centres. (Right) Kinetic fits of pyruvate and lactate signal from an automatically segmented tumour image ROI.

Figure 2: Spectral time series for slice 3 (top left), 2 (top right), 1 with default settings (bottom left), 1 with slice-specific frequency optimization (bottom right). Dotted vertical lines indicate frequencies at which images are reconstructed.

Figure 3: Images summed over the time series for (top left) pyruvate slice 3, (top right) pyruvate slice 2, (bottom left) pyruvate slice 1, (bottom right) lactate slice 2.